# ChemQC Lot Report

| ## Lot Information Lot Number:; 20251103J. Analysis Date:; 2025-12-12. | ## Test Summary Total Channels Analyzed:; 6. Channels Passed:; 0. Channels Failed:; 6. Final Lot Disposition:; FAIL. |
| --- | --- |

## Lot Summary Result

| Ink | Green Channel (G) | Red Channel (R) |
| --- | --- | --- |
| 51.9 | FAIL | FAIL |
| 33.39 | FAIL | FAIL |
| 51.9 LD | FAIL | FAIL |

## Report Generation

| Operator: | Raphael Pham |
| --- | --- |
| Timestamp: | 2026-01-05 13:34:56 |

## Manual Verification

| Operator Initials: ____________________________________________________ | Date: ____________________________________________________ |
| --- | --- |

# 51.9 Analysis

## Lot: 20251103J

## Green Channel Intensity; FAIL. SO2 Sensors: 0 passed, 6 failed (Total: 6) ## Results Summary Parameter; Specification; Value; Result; Batch Specification; Batch Value; Batch Result. SO2; > 10%; N/A; N/A; N/A; N/A; N/A. Performance; 95% of wells pass; 56.25% 52.0833% 42.7083% 30.2083% 33.3333% 41.6667%; FAIL FAIL FAIL FAIL FAIL FAIL; 80% of sensors; 10.0%; FAIL. Uniformity; 1% (CV of Gci); 0.996% 0.7499% 0.7691% 0.7457% 0.8364% 0.7868%; PASS PASS PASS PASS PASS PASS; 80% of sensors; 100.0%; PASS.

## Red Channel Intensity; FAIL. SO2 Sensors: 0 passed, 6 failed (Total: 6) ## Results Summary Parameter; Specification; Value; Result; Batch Specification; Batch Value; Batch Result. SO2; > 5.46%; N/A; N/A; N/A; N/A; N/A. Performance; 95% of wells pass; 56.25% 52.0833% 42.7083% 30.2083% 33.3333% 41.6667%; FAIL FAIL FAIL FAIL FAIL FAIL; 80% of sensors; 10.0%; FAIL. Uniformity; 1% (CV of Gci); 0.5486% 0.772% 0.4992% 0.4564% 0.5366% 0.7064%; PASS PASS PASS PASS PASS PASS; 80% of sensors; 100.0%; PASS.

# 33.39 Analysis

## Lot: 20251103J

## Green Channel Intensity; FAIL. NH3 Sensors: 0 passed, 4 failed (Total: 4) ## Results Summary Parameter; Specification; Value; Result; Batch Specification; Batch Value; Batch Result. NH3; < -20%; N/A; N/A; N/A; N/A; N/A. Performance; 95% of wells pass; 93.75% 94.7917% 92.7083% 95.8333%; FAIL FAIL FAIL PASS; 80% of sensors; 10.0%; FAIL. Uniformity; 1% (CV of Gci); 0.3472% 0.3734% 0.3713% 0.3917%; PASS PASS PASS PASS; 80% of sensors; 100.0%; PASS.

## Red Channel Intensity; FAIL. NH3 Sensors: 0 passed, 4 failed (Total: 4) ## Results Summary Parameter; Specification; Value; Result; Batch Specification; Batch Value; Batch Result. NH3; < -20%; N/A; N/A; N/A; N/A; N/A. Performance; 95% of wells pass; 93.75% 94.7917% 92.7083% 95.8333%; FAIL FAIL FAIL PASS; 80% of sensors; 10.0%; FAIL. Uniformity; 1% (CV of Gci); 0.349% 0.3359% 0.3501% 0.401%; PASS PASS PASS PASS; 80% of sensors; 100.0%; PASS.

# 51.9 LD Analysis

## Lot: 20251103J

## Green Channel Intensity; FAIL. SO2 Sensors: 0 passed, 6 failed (Total: 6) ## Results Summary Parameter; Specification; Value; Result; Batch Specification; Batch Value; Batch Result. SO2; > 10%; N/A; N/A; N/A; N/A; N/A. Performance; 95% of wells pass; 63.6364% 57.1429% 57.1429% 50.6494% 54.5455% 53.2468%; FAIL FAIL FAIL FAIL FAIL FAIL; 80% of sensors; 10.0%; FAIL.

## Red Channel Intensity; FAIL. SO2 Sensors: 0 passed, 6 failed (Total: 6) ## Results Summary Parameter; Specification; Value; Result; Batch Specification; Batch Value; Batch Result. SO2; > 5.49%; N/A; N/A; N/A; N/A; N/A. Performance; 95% of wells pass; 63.6364% 57.1429% 57.1429% 50.6494% 54.5455% 53.2468%; FAIL FAIL FAIL FAIL FAIL FAIL; 80% of sensors; 10.0%; FAIL.